## Slide 1: Device level thermal characterization for electronic and photonics devices
Dr. Mo Shakouri
Chairman
Microsanj, LLC., Silicon Valley USA
info@microsanj.com

## Slide 2: Applications

## Slide 3: Outline
Challenges on thermal characterization
  For electronic, photonic and power devices
Thermoreflectance method
  Method and equipment
Example of various cases
  Hot spot needs a spatial resolution
  Transient response (100s ms) tells the logic sequences
  Emission overlay gives location of electron collisions
  Heat diffuses through the metal layers
  Thermal information provide potential Thermal map changes before device fails
  HBT Thermal Image
  Non-uniform device turn on in a symmetrical Silicon Controlled Rectifier
Photonics and power devices
  LED - Single junction, encapsulation, multi junction
  Multi junction PV cell
  InP laser device (cross section)
  Power transistor array
Summary

## Slide 4: Challenges on thermal characterization
General challenges for electronics devices
Small features: 10s nm – 100s microns – difficult to contact
High speed response due to the small thermal mass
Highly non-uniform
Additional challenges for photonics and power devices
Light emission (photonics)
High heat density
Heat sinks requirement (power devices)

## Slide 5: Non-contact and transient thermal-imaging
10-9
10-6
10-12
FASTER
SLOWER
THERMAL TRANSIENT [Sec]
10-3
100
10
1-D hotspot
Single
Tr
Nano
wires
S.L. Cooler
Logic Chips
TE Modules
TIM
thin film
substrate level
Pkg.
IC arrays
ESD
DEVICE / TECHNOLOGY
10-9
10-6
SMALLER
LARGER
10-3
10-2
DIMENSIONS [m]
Motivations
Scaling on space and time
Moving hotspots – issue of scanning methods
Size of hot spots – issue on thermo couples/point sensors.

## Slide 6: Thermoreflectance imaging setup
Microscope setup
Console box
DUT
LED
CCD
Objective lens
Sig. gen.
Control box.
Temp. contl.

## Slide 7: How it works - thermoreflectance
System diagram
Thermoreflectance coefficient
LED driver
GP-IB
Power LED
PC
Detector
CCD, InGaAs
Pulse generator & power amp
Microscope
Objective
Device
Thermal bed
Light

## Slide 8: Lock-in signals
Timing chart
Temperature data point along the bias cycle
Acquisition timing (shifting by cycle)
4ms @ 25% Duty Cycle
1ms
33ms @ 30Hz
t0
t0
100ms
delay
Device Excitation
CCD exposure
LED pulse
Temperature
t1
t1

## Slide 9: Through silicon and emission
Top view
InGaAs CCD
1300 nm
LED
Objective
Substrate
Flip Chip DUT
Bottom view
(Image from http://www.janis.com)
Transmittance vs Wavelength, Si
% Transmittance
Wavelength, mm
1.0
10.0
resolution

## Slide 10: Defects and signature of potential failure
Transient irregular timing - potential of logic/operation failure
Thermal foot print  irregular local energy spot
Arrhenius's law
Emission – sign of high density of electron collisions
Thermal hotspot – location of potential long-term reliability
Near Infrared (NIR) wavelength provide a capability of both thermal signal and emission simultaneously.
LED options: 1050, 1200, 1300, and 1500 nm

## Slide 11: Resolution and sensitivity
Temperature
Spatial resolution
d ≈ l/2
Visible wavelengths, d ≈ 250-300 nm
NIR d ≈ 500 nm
Time resolution
Emission
InGaAs uncooled camera effective sensitivity of one pixel for emission ~ 30 mW/mm2
n : number of averaging
due to the weak signal (Cth ~ 10-4 order)
As scaling smaller, time resolution must be smaller due to thermal diffusion.
Dt : 100ns for our setup.
(for 1% error in temperature)

## Slide 12: Transient thermal/emission imaging
Method | Resolution |  |  | Imag-ing? | Notes
 | x(mm) | T (K) | t (sec) |  | 
m Thermocouple | 50 | 0.01 | 0.1-10 | No | Contact method
IR Thermography | 3-10 | 0.02-1 | 1m | Yes | Emissivity dependent
Lock-in Thermog. | 3-10 | 1m | NA | Yes | Need cycling
Liquid Crystal Thermography | 2-5 | 0.5 | 100 | Yes | Only near phase transition (aging issues)
Thermo-reflectance | 0.3-0.5 | 0.08 | 800p- 0.1m | Yes | Need cycling
Optical scanning Interferometry | 0.5 | 100m | 6n- 0.1m | Scan | Indirect measurement (expansion)
Micro Raman | 0.5 | 1 | 10n | Scan | 3D T-distribution
Scanning thermal microscopy (SThM) | 0.05 | 0.1 | 10-100m | Scan | Contact method surface morphology
Emission Microscopy (EMMI) | 0.25 s.im.lens | - | Op lock-in | Yes | Emitted Photon density

## Slide 13: Small hotspot
a)
b)
1.4 mm gate on MOSFET
1.4 mm gate on MOSFET
Distance (mm)
0
10
20
30
40
40
50
60
Temperature (a.u.)
2
4
6
8
10
12
0

## Slide 14: Transient Behavior of IC Latch-Up
- Potential timing failure -
0.5 ms
0.7 ms
0.9 ms
1.0 ms
3.0 ms
Movie1
The latch-up location is circled in yellow

## Slide 15: Thermal and emission overlay images
Through silicon substrate, 450 mm thick.
LED l = 1300nm and
InGaAs camera (640 x 512)
5x
50x
Thermal signals
Emission signals
44 mW

## Slide 16: Diffusion time/depth estimations
a: thermal diffusivity [m2/s]
m: depth of heat source
t: time to reach observing surface

## Slide 17: Through silicon, deep under the 6th metal layer
2.0 msec
0.97V, ~12mA, ~12mW 20% duty cycle
10 minutes of averaging (repeating)
Movie2

## Slide 18: Time delay to reach to the surface
Precise time resolution is a key to find the response.

## Slide 19: Via chain reliability
Past work on reliability analysis – oven test for f500nm copper via through silicon.
K. Yazawa, D. Kendig, D. Hernandez, K. Maize, S. Alavi, A. Shakouri, “High Speed Transient Thermoreflectance Imaging of Microelectronic Devices and Circuits”, EDFA Magazine, Vol. 15, No 1. Page 12-22, (2013).
Rvia(t)/Rvia(0)
DTvia(t)/DTvia(0)

## Slide 20: Power Amplifiers
Copyright 2013 Microsanj, LLC All rights reserved.

## Slide 21: HBT Thermal Image
CCD Image
Thermal Image
20x
0.5Hz, 3.8W
380mA, 10V
∆T

## Slide 22: pHEMT Gate Temperature
CCD Image
Thermal Image
This shows a low magnification thermoreflectance image of pHEMT
20x

## Slide 23: 100x thermal image of gate
This shows the side temperature profile of the center of the pHEMT. The dark channel region is not calibrated for, and did not show a strong signal.
(calibrated for metal only)
CCD Image
Thermal Image

## Slide 24: 100x Side Profile
Y-Z Profile
44mA, 1V
This shows the side temperature profile of the center of the pHEMT. The dark channel region is not calibrated for, and did not show a strong signal.
(calibrated for metal only)
24

## Slide 25: Transient 3D Temperature Profiles
50us temperature map
500us temperature map
These images show the heat spreading from the gate/channel region to the source/drain region at 50µs and 500µs.

## Slide 26: ESD Protection Devices
Copyright 2013 Microsanj, LLC All rights reserved.

## Slide 27: 100 ns temporal resolution of ESD-type event
10µs
30µs
70µs
200µs
∆T (˚C)
ΔT [K]
20
15
10
5
0
Anode
Cathode
25µm
(a) SCR optical image
(b) SCR thermoreflectance in snapback after 300ns at 1.28A
Non-uniform device turn on in a symmetrical Silicon Controlled Rectifier (SCR) in response to a 2.5 ms pulse
Initial response to the 300 ns 120V pulse shows highly localized heating.
50µm
300ns
K. Maize, V. Vashchenko et al, IRPS, 2011.

## Slide 28: SCR Thermoreflectance Images
Optical and Thermal Image Overlay
100x Optical Image

## Slide 29: Transient SCR response
-These thermal images show the transient thermal response of the SCR.
-The SCR starts to conduct at the corners of the Anode (due to high electric field)
-Images are all on the same temperature scale
Scaled 100µs image to
Show temperature distribution
10V was applied to the device. The current and precise voltage across the device was not measured for this series
29

## Slide 30: Transient SCR response
High Electric Field at sharp corners

## Slide 31: GGNFET Transient Images
- These are transient thermoreflectance images at 1, 10, & 100 µs
- Each image is scaled to show the temperature distribution.
31

## Slide 32: GGNFET, 200ns, 30V pulse
Chart: Time ( µs )
| 0.0 | 0.05 | 0.1 | 0.15 | 0.2 | 0.25 | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 | 0.65 | 0.7 | 0.75 | 0.75 | 1.0 | 1.5 | 2.0 | 2.5 | 3.0 | 4.0 | 5.0 | 6.0 | 7.0 | 8.0 | 9.0 | 10.0 | 15.0
 | 0.26168 | -0.37937 | 0.60882 | 0.93108 | 3.4212 | 6.9653 | 9.6773 | 12.0757 | 12.796 | 14.8118 | 15.9666 | 17.472 | 18.0149 | 17.8288 | 18.2777 | 18.1092 | 17.9176 | 16.5754 | 12.3856 | 9.6917 | 7.4939 | 6.3933 | 4.1966 | 3.2324 | 2.437 | 2.1201 | 1.9074 | 1.5671 | 1.3 | 0.5
Chart: Time ( µs )
| 0.0 | 0.05 | 0.1 | 0.15 | 0.2 | 0.25 | 0.3 | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 | 0.65 | 0.7 | 0.75 | 0.75 | 1.0 | 1.5
 | 0.26168 | -0.37937 | 0.60882 | 0.93108 | 3.4212 | 6.9653 | 9.6773 | 12.0757 | 12.796 | 14.8118 | 15.9666 | 17.472 | 18.0149 | 17.8288 | 18.2777 | 18.1092 | 17.9176 | 16.5754 | 12.3856
This shows the temperature rise (a.u.) of the GGNFET in response to a 200 ns 30V pulse

## Slide 33: Photonics and Power device examples
LEDs
  single junction, encapsulation, multi junction
Multi junction PV cell
InP laser device (cross section)
Power transistor array

## Slide 34
D. Kendig et al, SemiTHERM, 2011
Separating electroluminescence and thermal information

## Slide 35: Thermal imaging of an encapsulated LED
Non-calibrated thermal image
Electroluminescence Image
10x

## Slide 36
a
a
a’
a’
a
a
a’
a’
LED2
LED1
There seems to be more heating closer to the electrodes
LED
peak 447.5nm
Transient thermoreflectance Imaging on LED

## Slide 37: Temperature profiles – another section
a
a’
a
a’
In the Y direction profile, the heating is more uniform.

## Slide 38: Multi-junction PV
125 microns
Emission at forward bias 3V, 140mA
10 microns
250x
20x
1.25x
optical
emission
20x
=
http://www1.eere.energy.gov/solar/sunshot/pv_high_efficiency.html
DoE SunShot initiative
This is hard to find by
optical inspection only
Device structure

## Slide 39: InP Laser device (cross section) with 530nm and 470nm LEDs
Forward diode voltage 0.73V
Chart: Current (A)
| 0.075 | 0.1 | 0.125 | 0.15 | 0.175 | 0.2 | 0.225
530nm | 2.0955 | 3.6482 | 5.1844 | 6.6993 | 8.8208 | 11.1089 | 13.3819
470nm | 2.76 | 4.25 | 5.99 | 7.94 | 10.3 | 12.7 | 15.8
Average | 2.4278 | 3.9491 | 5.5872 | 7.3197 | 9.5604 | 11.9045 | 14.591
Laser Ridge | 3.1678 | 5.9906 | 9.138 | 9.9245 | 11.5109 | 16.069 | 17.2734
Areas of interest (0.5 mm x 1 mm)
In flip chip package
250x

## Slide 40
100x
250mA transient thermal images at 10, 20, 100, 200ms

## Slide 41: Transient response (250mA pulse)
100x, 10us
A
B
Chart: 
| 1.0E-6 | 2.0E-6 | 3.0E-6 | 4.0E-6 | 5.0E-6 | 6.0E-6 | 7.0E-6 | 8.0E-6 | 9.0E-6 | 1.0E-5 | 1.1E-5 | 1.2E-5 | 1.3E-5 | 1.4E-5 | 1.5E-5 | 1.6E-5 | 1.7E-5 | 1.8E-5 | 1.9E-5 | 2.0E-5 | 2.5E-5 | 3.0E-5 | 3.5E-5 | 4.0E-5 | 4.5E-5 | 5.0E-5 | 5.5E-5 | 6.0E-5 | 6.5E-5 | 7.0E-5 | 7.5E-5 | 8.0E-5 | 8.5E-5 | 9.0E-5 | 9.5E-5 | 0.0001 | 0.000105 | 0.00011
InPQWsm | 0.202 | 2.74 | 8.62 | 10.6 | 14.3 | 14.1 | 14.4 | 14.1 | 14.5 | 14.1 | 14.9 | 15.1 | 15 | 16.4 | 14.8 | 15.2 | 14.9 | 15.6 | 15.7 | 15.4 | 15.7 | 15.8 | 15.8 | 16.3 | 16.6 | 15.9 | 5.55 | 3.5 | 2.88 | 2.24 | 1.57 | 1.91 | 1.68 | 1.04 | 1.14 | 1.48 | 1.05 | 0.153
InPsm | -0.0419 | 0.0677 | 0.103 | 0.323 | 0.744 | 1.09 | 1.4 | 1.83 | 2.02 | 2.3 | 2.58 | 2.74 | 2.97 | 3.1 | 3.29 | 3.53 | 3.65 | 3.8 | 3.93 | 4.05 | 4.54 | 4.83 | 5.13 | 5.42 | 5.82 | 5.83 | 4.88 | 3.55 | 2.73 | 2.25 | 1.91 | 1.51 | 1.36 | 1.19 | 1.03 | 0.916 | 0.704 | 0.72
InPxsm | -0.101 | 0.338 | 0.432 | 0.892 | 1.6 | 2.11 | 2.38 | 3.08 | 3.4 | 3.7 | 4.21 | 4.04 | 4.71 | 4.63 | 4.98 | 5.13 | 5.29 | 5.32 | 5.45 | 5.76 | 6.51 | 6.64 | 7.15 | 7.27 | 7.78 | 7.73 | 6 | 3.93 | 3.1 | 2.52 | 2.17 | 1.86 | 1.37 | 1.27 | 1.37 | 1.04 | 0.638 | 0.88
Time [sec]
Temperature change [K]
A: InP substrate
B: InP based modified material

## Slide 42: Thermoreflectance coefficient characterization for new materials
Chart: Wave length of LED [nm]
| 470.0 | 530.0 | 780.0
 | 0.000188 | 0.000156 | 6.33e-05
 | 0.00019149 | 0.0001619 | 8.43e-05
 | 0.00018451 | 0.0001501 | 4.23e-05
InP: Fe doped

## Slide 43: High Power MOSFET Transistor Array
“Thermal Characterization of High Power Transistor Arrays”, K. Maize et al, 25th IEEE SEMI-THERM Symposium, 2009.

## Slide 44: Thermal image vs. simulation at low bias
b
0.25
0.3
0.35
0.4
0.45
0.5
0.55
0.6
0.65
a
a’
The magnitudes of temperature rise at both source and drain are in a good agreement.
Vd = 0.5V, Id ≈ 0.5A 6 amps/mm2

## Slide 45: Array heating at low & high current densities
Δ Temperature [°C]
47 A/mm2
VD=5V, ID=4.2A
b
b’
Δ T profile from b to b'
Distance
30
70
50
60
40
ΔT [°C]
HIGH BIAS
60
20
40
0
Δ Temperature [°C]
b
b’
400μm
6 A/mm2
VD=500mV, ID=545mA
a
a’
Distance
LOW BIAS
400μm
a
a’
0.4
0.6
0.2
0
ΔT [°C]
Δ Temperature [°C]
0.4
0.6
0.2
0.5
0.7
0.3
Δ T profile from a to a'
Thermo-reflectance thermal images
Tempera-ture profile along array
“Thermal Characterization of High Power Transistor Arrays”, K. Maize et al, 25th IEEE SEMI-THERM Symposium, 2009.

## Slide 46: Comparison of topside and backside heating
Near-IR Light
Thermal images
Topside
Backside
(InGaAs CCD)
drain contact
source contact
100µm
6X
K. Maize et al, SemiTHERM 2010 and IHTC 2010.

## Slide 47: Microsanj Products

## Slide 48: Key System Specifications
Product specifications & configurations are subject to change.
1. Requires optical table mount

## Slide 50: Publications
Bias-Dependent MOS Transistor Thermal Resistance and Non-Uniform Self-heating Temperature
Si/Ge uni-traveling carrier photodetector
High-power high-linearity flip-chip bonded modified uni-traveling carrier photodiode
Time and Frequency Domain CCD-Based Thermoreflectance Techniques for High-Resolution Transient Thermal Imaging
Understanding the Thermoreflectance Coefficient for High Resolution Thermal Imaging of Microelectronic Devices
High Speed Transient Thermal Imaging of Microelectronic Devices and Circuits
Thermal Characterization of High Power Transistor Arrays
Time and Frequency Domain CCD-Based Thermoreflectance Techniques for High-Resolution Transient Thermal Imaging
Thermoreflectance Imaging of Defects in Thin-Film Solar Cells
Thermal Imaging of Encapsulated LEDs
Quantum Electronics Group, Jack Baskin School of Engineering at UC Santa Cruz
Thermal Imaging for Reliability Characterization of Copper Vias
Ultrafast Submicron Thermoreflectance Imaging
More information at http://microsanj.com/news-and-papers
50

## Slide 51: Summary
Introduced thermoreflectance imaging corresponding to challenges on thermal characterization of electronics devices.
Discussed the transient thermal characterization in 2D images in sub-micron length scale.
Demonstrated the thermal characterization of photonic devices and power devices.

## Slide 52
Microsanj社の 開発した熱画像解析装置、Nanothermシリーズは、これまでのIRによるサーモグラフィー装置とは 全く異なった温度測定技術を用いたシステムです。測定物のIR放射を測定するのではなく、 測定物に非常に短時間の光を照射し、その反射光を計測することにより温度分布を測定するため、 測定物に全く影響を与えること無く、IRでは難しかった広い温度範囲を非接触にて測定することが 可能となりました。測定は金属を含むあらゆるものが可能で、測定物を熱したり、表面に特別な処理を 行う必要が有りません。また、薄いシリコン基板は光を透過することから、flip-chip等の、シリコン基板上の 半導体の熱画像を裏面から観測することが可能です。また、Nanothermシステムの最大の特徴として、 オプションにてバイアス電源と信号源を追加することにより、熱画像の過渡特性を、最速では0.8nsec間隔で 測定することができます。Nanothermシステムにより、温度上昇、熱集中の状況をリアルタイムに観察することで、 半導体そのものや半導体回路の最適な熱設計を行うこと、また故障解析、不良解析を行うことが可能です。 測定物の大きさは最小300nm、温度分解能は最小0.2℃、測定温度範囲-265～500℃に対応します。
ATN Japan
1-35-16 Nakagawa-Chuo
Tsuzuki, Yokohama, Kanagawa, 224-0003 JAPAN
Website: www.atnjapan.com
E-mail: sales@atnjapan.com